## Savings Summary
Savings Calculation:
|  | ULT Freezer Range: 15 to <24 ft3 (- 75 °C) | [6 empty] | ULT Freezer Range: 15 to <24 ft3 (- 75 °C) | [8 empty] | ULT Freezer Range: 24 to 29 ft3 (- 75 °C)
|  | Average Volume | Direct HE ULT Freezer Savings | Direct HE ULT Freezer Savings | Direct HE ULT Freezer Demand Reduction | Indirect HVAC Savings | Indirect HVAC Demand Reduction | CZ | Indirect HVAC Gas Penalty | Total HE ULT Freezer Savings | Total HE ULT Freezer Demand Reduction | Average Volume | Direct HE ULT Freezer Savings | Direct HE ULT Freezer Savings | Direct HE ULT Freezer Demand Reduction | Indirect HVAC Savings | Indirect HVAC Demand Reduction | CZ | Indirect HVAC Gas Penalty | Total HE ULT Freezer Savings | Total HE ULT Freezer Demand Reduction
|  | (ft3) | (kWh/ft3-day) | (kWh/yr-each) | (kW/each) | (kWh/yr-each) | (kW/each) |  | (therm/yr-each) | (kWh/yr-each) | (kW/each) | (ft3) | (kWh/ft3-day) | (kWh/yr-each) | (kW/each) | (kWh/yr-each) | (kW/each) |  | (therm/yr-each) | (kWh/yr-each) | (kW/each)
| CZ01 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 222.52 | 0.03034 | CZ01 | -14.8056 | 2082.649 | 0.24268 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 543 | 0.073008 | CZ01 | -34.8856 | 4282.4606 | 0.49989
| CZ02 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 263.54 | 0.010932 | CZ02 | -13.7476 | 2123.669 | 0.22328 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 606.86 | 0.10311 | CZ02 | -32.5208 | 4346.3206 | 0.52999
| CZ03 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 226.02 | 0.034886 | CZ03 | -15.488 | 2086.149 | 0.24723 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 566 | 0.08306 | CZ03 | -37.0562 | 4305.4606 | 0.50994
| CZ04 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 246.28 | 0.051394 | CZ04 | -15.335 | 2106.409 | 0.26374 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 597.98 | 0.12433 | CZ04 | -36.6274 | 4337.4406 | 0.55121
| CZ05 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 239.9 | 0.043112 | CZ05 | -14.8814 | 2100.029 | 0.25546 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 575.48 | 0.10329 | CZ05 | -36.0712 | 4314.9406 | 0.53017
| CZ06 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 244.14 | 0.043704 | CZ06 | -16.0476 | 2104.269 | 0.25605 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 685.06 | 0.1046 | CZ06 | -38.5444 | 4424.5206 | 0.53148
| CZ07 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 229.64 | 0.011684 | CZ07 | -16.6074 | 2089.769 | 0.22403 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 556.08 | 0.028084 | CZ07 | -40.2886 | 4295.5406 | 0.45496
| CZ08 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 212.5 | 0.016656 | CZ08 | -16.4444 | 2072.629 | 0.229 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 583.26 | 0.03973 | CZ08 | -38.9296 | 4322.7206 | 0.46661
| CZ09 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 276.86 | 0.051452 | CZ09 | -15.3958 | 2136.989 | 0.2638 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 698.34 | 0.12732 | CZ09 | -36.7108 | 4437.8006 | 0.5542
| CZ10 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 276.14 | 0.015594 | CZ10 | -15.0614 | 2136.269 | 0.22794 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 649.14 | 0.037798 | CZ10 | -35.9128 | 4388.6006 | 0.46468
| CZ11 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 268.78 | 0.045458 | CZ11 | -14.114 | 2128.909 | 0.2578 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 650.92 | 0.1096 | CZ11 | -34.0282 | 4390.3806 | 0.53648
| CZ12 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 241.5 | 0.049084 | CZ12 | -14.7542 | 2101.629 | 0.26143 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 582.38 | 0.11789 | CZ12 | -35.5028 | 4321.8406 | 0.54477
| CZ13 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 267.3 | 0.046308 | CZ13 | -14.579 | 2127.429 | 0.25865 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 645.56 | 0.11103 | CZ13 | -35.1134 | 4385.0206 | 0.53791
| CZ14 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 285.7 | 0.020234 | CZ14 | -11.4754 | 2145.829 | 0.23258 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 651 | 0.048494 | CZ14 | -28.5598 | 4390.4606 | 0.47537
| CZ15 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 336.42 | 0.005232 | CZ15 | -16.566 | 2196.549 | 0.21758 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 740.2 | 0.012926 | CZ15 | -39.449 | 4479.6606 | 0.43981
| CZ16 | 19.5 | 0.26135 | 1860.129 | 0.21234 | 255.08 | 0.01822 | CZ16 | -7.9828 | 2115.209 | 0.23056 | 26.5 | 0.38661 | 3739.4606 | 0.42688 | 629.8 | 0.043114 | CZ16 | -18.662 | 4369.2606 | 0.46999
[8 empty] | Average | -14.5804 | 2115.899 | 0.24324 | [6 empty] | Average | -34.9289 | 4362.0269 | 0.50609
|  | Base Model: | 500 | Watt ULT freezer per space type | [5 empty] | 38.7% | 38.97% | Base Model: | 750 | Watt ULT freezer per space type
|  | Proposed Model: | 329.9088 | Watt ULT freezer per space type | [5 empty] | 5468.0338 | 0.6242 | Proposed Model: | 341.7973 | Watt ULT freezer per space type
3607.9048
1860.129
Reference: | ET, Tables 6, 10 and 11, Figure 27
|  |  | -80 |  | -70 |  | -75
ULT Freezer | Volume | Refrigerant | EC, -80C sp |  | EC, -70C sp |  | EC, -75C sp, est |  |  |  | Marked as
| (ft3) | Type | (kWh/day) | (kWh/ft3/day) | (kWh/day) | (kWh/ft3/day) | (kWh/day) | (kWh/ft3/day) | (kWh/day) | (kWh/ft3/day) | Energy Efficent?
K | 16 | HFC | 14.66 | 0.91625 | 11.06 | 0.69125 | 12.94 | 0.80875 | 12.78 | 0.79875 |  |  | 12.86 | 0.80375
M | 18 | HFC | 14.84 | 0.82444 | 11.72 | 0.65111 | 13.22 | 0.73444 | 12.74 | 0.70778 |  |  | 13.28 | 0.73778
N | 18.9 | HFC | 20.43 | 1.081 | 14.99 | 0.79312 | 17.82 | 0.94286 | 17.01 | 0.9 |  |  | 17.71 | 0.93704
H | 19.4 | Natural | 10.44 | 0.53814 | 7.06 | 0.36392 | 9.08 | 0.46804 | 8.35 | 0.43041 | Yes |  | 8.75 | 0.45103
A | 20.1 | HFC | 12.64 | 0.62886 | 8.68 | 0.43184 | 10.97 | 0.54577 | 10.04 | 0.4995 | Yes |  | 10.66 | 0.53035
B | 23 | HFC/Natural Blend | 14.89 | 0.64739 | 11.78 | 0.51217 | 13.5 | 0.58696 | 12.65 | 0.55 |  |  | 13.335 | 0.57978
Ave. EE |  |  |  | 0.5835 |  | 0.39788 |  | 0.50691 |  | 0.46496 |  |  |  | 0.49069
Ave. SE |  |  |  | 0.86726 |  | 0.66191 |  | 0.76825 |  | 0.73913 |  |  |  | 0.76459
Savings |  |  |  | 0.28376 |  | 0.26404 |  | 0.26135 |  | 0.27417 |  |  |  | 0.2739
|  |  |  | 32.72% |  | 39.89% |  | 34.02%
Average | 19.5 | (ft3) |  |  |  |  | 598.75 | watt/day | 0.34018
Savings | 1860.129 | kWh/yr-case |  | 0.21234
D | 24 | HFC | 20.46 | 0.8525 | 14.04 | 0.585 | 17.88 | 0.745 | 17.1 | 0.7125 |  |  | 17.25 | 0.71875
C | 24.7 | HFC | 22.52 | 0.91174 | 15.11 | 0.61174 | 19.03 | 0.77045 | 17.81 | 0.72105 |  |  | 18.815 | 0.76174
E | 25.7 | HFC | 26.71 | 1.0393 | 19.85 | 0.77237 | 23.76 | 0.92451 | 22.52 | 0.87626 |  |  | 23.28 | 0.90584
L | 25.7 | HFC | 18.89 | 0.73502 | 14.11 | 0.54903 | 16.74 | 0.65136 | 15.86 | 0.61712 |  |  | 16.5 | 0.64202
O | 26 | HFC | 17.72 | 0.68154 | 12.64 | 0.48615 | 15.17 | 0.58346 | 14.66 | 0.56385 |  |  | 15.18 | 0.58385
F | 27.5 | Natural | 9.68 | 0.352 | 7.89 | 0.28691 | 8.83 | 0.32109 | 7.72 | 0.28073 | Yes |  | 8.785 | 0.31945
I | 27.5 | Natural | 8.72 | 0.31709 | 6.96 | 0.25309 | 7.86 | 0.28582 | 6.86 | 0.24945 | Yes |  | 7.84 | 0.28509
G | 28.8 | Natural | 12.12 | 0.42083 | 8.05 | 0.27951 | 10.49 | 0.36424 | 9.44 | 0.32778 | Yes |  | 10.085 | 0.35017
J | 28.8 | HFC/Natural Blend | 19.96 | 0.69306 | 13.76 | 0.47778 | 16.91 | 0.58715 | 15.35 | 0.53299 |  |  | 16.86 | 0.58542
Ave. EE |  |  |  | 0.36331 |  | 0.27317 |  | 0.32372 |  | 0.28599 |  |  |  | 0.31824
Ave. SE |  |  |  | 0.81886 |  | 0.58035 |  | 0.71032 |  | 0.67063 |  |  |  | 0.6996
Savings |  |  |  | 0.45555 |  | 0.30717 |  | 0.38661 |  | 0.38464 |  |  |  | 0.38136
|  |  |  | 55.63% |  | 52.93% |  | 54.43%
Average | 26.5 | (ft3) |  |  |  |  | 760.3472 | watt/day | 0.54427
Savings | 3739.4606 | kWh/yr-case
Note on D23: Author: Estimated based on average tested power of standard efficient units at -75C
Note on N23: Author: Estimated based on average tested power of standard efficient units at -75C
Note on H32: Author: This is with door openings, Table 11.
Note on J32: Author: This is with no door openings, Table 9.

## CZ01
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 182.7513 | 182.5813 | 182.5813 | 182.5813 | 182.5813
| Total | 3347448 | 3345959 | 3345657 | 3345920 | 3345489 |  | Total | 456.6647 | 456.4952 | 456.433 | 456.4909 | 456.4742
Natural Gas Consumption (kBtu) | Total | 9845826 | 9851542 | 9845662 | 9846340 | 9845834
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1489 | 1791 | 1528 | 1959 |  | Total |  | 0.1695 | 0.2317 | 0.1738 | 0.1905
Natural Gas Savings (therms) | Total |  | -57.16 | 1.64 | -5.14 | -0.08
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0005 | 0.0617 | 0.0038 | 0.0205
| Indirect HVAC Savings, kWh |  | -3 | 299 | 36 | 467 |  | Weight Reduction, kW |  |  |  |  | 0.03034 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 222.52 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -57.16 | 1.64 | -5.14 | -0.08
| Weight Savings, Therm |  |  |  |  | -14.8056
ULT Freezer Range: 24 to 29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 183.7513 | 183.3433 | 183.3433 | 183.3437 | 183.3437
| Total | 3357404 | 3353813 | 3353070 | 3353723 | 3352691 |  | Total | 457.7906 | 457.3833 | 457.2349 | 457.3732 | 457.3331
Natural Gas Consumption (kBtu) | Total | 9837110 | 9850670 | 9836662 | 9838336 | 9837130
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.4076 | 0.4076
| Total |  | 3591 | 4334 | 3681 | 4713 |  | Total |  | 0.4073 | 0.5557 | 0.4174 | 0.4575
Natural Gas Savings (therms) | Total |  | -135.6 | 4.48 | -12.26 | -0.2
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0007 | 0.1477 | 0.0098 | 0.0499
| Indirect HVAC Savings, kWh |  | -5 | 739 | 86 | 1118 |  | Weight Reduction, kW |  |  |  |  | 0.073008 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 543 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -135.6 | 4.48 | -12.26 | -0.2
| Weight Savings, Therm |  |  |  |  | -34.8856

## CZ02
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 182.7513 | 182.5813 | 182.5813 | 182.5813 | 182.5813
| Total | 3472333 | 3470838 | 3470427 | 3470802 | 3470420 |  | Total | 510.3714 | 510.202 | 510.19 | 510.1959 | 510.1732
Natural Gas Consumption (kBtu) | Total | 9615558 | 9621207 | 9615185 | 9616067 | 9615566
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1495 | 1906 | 1531 | 1913 |  | Total |  | 0.1694 | 0.1814 | 0.1755 | 0.1982
Natural Gas Savings (therms) | Total |  | -56.49 | 3.73 | -5.09 | -0.08
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0006 | 0.0114 | 0.0055 | 0.0282
| Indirect HVAC Savings, kWh |  | 3 | 414 | 39 | 421 |  | Weight Reduction, kW |  |  |  |  | 0.010932 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 263.54 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -56.49 | 3.73 | -5.09 | -0.08
| Weight Savings, Therm |  |  |  |  | -13.7476
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 183.7513 | 183.3433 | 183.3433 | 183.3437 | 183.3437
| Total | 3482347 | 3478746 | 3477818 | 3478658 | 3477742 |  | Total | 511.5491 | 511.1416 | 510.9319 | 511.1274 | 511.0731
Natural Gas Consumption (kBtu) | Total | 9607018 | 9620557 | 9606017 | 9608256 | 9607037
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.4076 | 0.4076
| Total |  | 3601 | 4529 | 3689 | 4605 |  | Total |  | 0.4075 | 0.6172 | 0.4217 | 0.476
Natural Gas Savings (therms) | Total |  | -135.39 | 10.01 | -12.38 | -0.19
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0005 | 0.2092 | 0.0141 | 0.0684
| Indirect HVAC Savings, kWh |  | 5 | 934 | 94 | 1010 |  | Weight Reduction, kW |  |  |  |  | 0.10311 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 606.86 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -135.39 | 10.01 | -12.38 | -0.19
| Weight Savings, Therm |  |  |  |  | -32.5208

## CZ03
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 182.7513 | 182.5813 | 182.5813 | 182.5813 | 182.5813
| Total | 3434209 | 3432720 | 3432377 | 3432682 | 3432318 |  | Total | 500.1312 | 499.9614 | 499.8913 | 499.9567 | 499.936
Natural Gas Consumption (kBtu) | Total | 8228603 | 8234270 | 8228632 | 8229118 | 8228610
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1489 | 1832 | 1527 | 1891 |  | Total |  | 0.1698 | 0.2399 | 0.1745 | 0.1952
Natural Gas Savings (therms) | Total |  | -56.67 | -0.29 | -5.15 | -0.07
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0002 | 0.0699 | 0.0045 | 0.0252
| Indirect HVAC Savings, kWh |  | -3 | 340 | 35 | 399 |  | Weight Reduction, kW |  |  |  |  | 0.034886 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 226.02 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -56.67 | -0.29 | -5.15 | -0.07
| Weight Savings, Therm |  |  |  |  | -15.488
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 183.7513 | 183.3433 | 183.3433 | 183.3437 | 183.3437
| Total | 3444144 | 3440553 | 3439692 | 3440466 | 3439562 |  | Total | 501.2757 | 500.8688 | 500.701 | 500.857 | 500.8081
Natural Gas Consumption (kBtu) | Total | 8219482 | 8233100 | 8219513 | 8220718 | 8219500
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.4076 | 0.4076
| Total |  | 3591 | 4452 | 3678 | 4582 |  | Total |  | 0.4069 | 0.5747 | 0.4187 | 0.4676
Natural Gas Savings (therms) | Total |  | -136.18 | -0.31 | -12.36 | -0.18
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0011 | 0.1667 | 0.0111 | 0.06
| Indirect HVAC Savings, kWh |  | -5 | 857 | 83 | 987 |  | Weight Reduction, kW |  |  |  |  | 0.08306 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 566 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -136.18 | -0.31 | -12.36 | -0.18
| Weight Savings, Therm |  |  |  |  | -37.0562

## CZ04
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 198.129 | 197.959 | 197.959 | 197.959 | 197.959
| Total | 3590350 | 3588859 | 3588467 | 3588822 | 3588468 |  | Total | 648.571 | 648.4021 | 648.2934 | 648.3946 | 648.3724
Natural Gas Consumption (kBtu) | Total | 8573454 | 8579140 | 8573433 | 8573972 | 8573463
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1491 | 1883 | 1528 | 1882 |  | Total |  | 0.1689 | 0.2776 | 0.1764 | 0.1986
Natural Gas Savings (therms) | Total |  | -56.86 | 0.21 | -5.18 | -0.09
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0011 | 0.1076 | 0.0064 | 0.0286
| Indirect HVAC Savings, kWh |  | -1 | 391 | 36 | 390 |  | Weight Reduction, kW |  |  |  |  | 0.051394 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 246.28 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -56.86 | 0.21 | -5.18 | -0.09
| Weight Savings, Therm |  |  |  |  | -15.335
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 199.129 | 198.721 | 198.721 | 198.721 | 198.721
| Total | 3600326 | 3596733 | 3595777 | 3596642 | 3595794 |  | Total | 649.7814 | 649.3739 | 649.1146 | 649.3574 | 649.3042
Natural Gas Consumption (kBtu) | Total | 8564364 | 8577989 | 8564284 | 8565596 | 8564394
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.408 | 0.408
| Total |  | 3593 | 4549 | 3684 | 4532 |  | Total |  | 0.4075 | 0.6668 | 0.424 | 0.4772
Natural Gas Savings (therms) | Total |  | -136.25 | 0.8 | -12.32 | -0.3
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0005 | 0.2588 | 0.016 | 0.0692
| Indirect HVAC Savings, kWh |  | -3 | 954 | 89 | 937 |  | Weight Reduction, kW |  |  |  |  | 0.12433 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 597.98 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -136.25 | 0.8 | -12.32 | -0.3
| Weight Savings, Therm |  |  |  |  | -36.6274

## CZ05
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 198.129 | 197.959 | 197.959 | 197.959 | 197.959
| Total | 3452651 | 3451161 | 3450786 | 3451121 | 3450763 |  | Total | 536.1391 | 535.9693 | 535.8809 | 535.9636 | 535.9418
Natural Gas Consumption (kBtu) | Total | 8375267 | 8380835 | 8375212 | 8375785 | 8375274
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1490 | 1865 | 1530 | 1888 |  | Total |  | 0.1698 | 0.2582 | 0.1755 | 0.1973
Natural Gas Savings (therms) | Total |  | -55.68 | 0.55 | -5.18 | -0.07
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0002 | 0.0882 | 0.0055 | 0.0273
| Indirect HVAC Savings, kWh |  | -2 | 373 | 38 | 396 |  | Weight Reduction, kW |  |  |  |  | 0.043112 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 239.9 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -55.68 | 0.55 | -5.18 | -0.07
| Weight Savings, Therm |  |  |  |  | -14.8814
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 199.129 | 198.721 | 198.721 | 198.721 | 198.721
| Total | 3462592 | 3459001 | 3458100 | 3458910 | 3458049 |  | Total | 537.3162 | 536.9091 | 536.6964 | 536.8948 | 536.8434
Natural Gas Consumption (kBtu) | Total | 8366236 | 8379760 | 8366088 | 8367483 | 8366253
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.408 | 0.408
| Total |  | 3591 | 4492 | 3682 | 4543 |  | Total |  | 0.4071 | 0.6198 | 0.4214 | 0.4728
Natural Gas Savings (therms) | Total |  | -135.24 | 1.48 | -12.47 | -0.17
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0009 | 0.2118 | 0.0134 | 0.0648
| Indirect HVAC Savings, kWh |  | -5 | 897 | 87 | 948 |  | Weight Reduction, kW |  |  |  |  | 0.10329 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 575.48 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -135.24 | 1.48 | -12.47 | -0.17
| Weight Savings, Therm |  |  |  |  | -36.0712

## CZ06
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 198.129 | 197.959 | 197.959 | 197.959 | 197.959
| Total | 3564469 | 3562979 | 3562560 | 3562941 | 3562651 |  | Total | 562.2886 | 562.119 | 562.0294 | 562.1129 | 562.0903
Natural Gas Consumption (kBtu) | Total | 6593430 | 6599090 | 6593592 | 6593961 | 6593437
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1490 | 1909 | 1528 | 1818 |  | Total |  | 0.1696 | 0.2592 | 0.1757 | 0.1983
Natural Gas Savings (therms) | Total |  | -56.6 | -1.62 | -5.31 | -0.07
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0004 | 0.0892 | 0.0057 | 0.0283
| Indirect HVAC Savings, kWh |  | -2 | 417 | 36 | 326 |  | Weight Reduction, kW |  |  |  |  | 0.043704 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 244.14 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -56.6 | -1.62 | -5.31 | -0.07
| Weight Savings, Therm |  |  |  |  | -16.0476
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 199.129 | 198.721 | 198.721 | 198.721 | 198.721
| Total | 3574570 | 3570980 | 3569800 | 3570859 | 3570079 |  | Total | 563.4684 | 563.0612 | 562.8468 | 563.0469 | 562.993
Natural Gas Consumption (kBtu) | Total | 6584065 | 6597658 | 6584456 | 6585332 | 6584085
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.408 | 0.408
| Total |  | 3590 | 4770 | 3711 | 4491 |  | Total |  | 0.4072 | 0.6216 | 0.4215 | 0.4754
Natural Gas Savings (therms) | Total |  | -135.93 | -3.91 | -12.67 | -0.2
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0008 | 0.2136 | 0.0135 | 0.0674
| Indirect HVAC Savings, kWh |  | -6 | 1175 | 116 | 896 |  | Weight Reduction, kW |  |  |  |  | 0.1046 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 685.06 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -135.93 | -3.91 | -12.67 | -0.2
| Weight Savings, Therm |  |  |  |  | -38.5444

## CZ07
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 198.129 | 197.959 | 197.959 | 197.959 | 197.959
| Total | 3613322 | 3611828 | 3611436 | 3611798 | 3611531 |  | Total | 600.0332 | 599.8608 | 599.8522 | 599.8612 | 599.8322
Natural Gas Consumption (kBtu) | Total | 6418852 | 6424464 | 6419147 | 6419408 | 6418907
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1494 | 1886 | 1524 | 1791 |  | Total |  | 0.1724 | 0.181 | 0.172 | 0.201
Natural Gas Savings (therms) | Total |  | -56.12 | -2.95 | -5.56 | -0.55
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0024 | 0.011 | 0.002 | 0.031
| Indirect HVAC Savings, kWh |  | 2 | 394 | 32 | 299 |  | Weight Reduction, kW |  |  |  |  | 0.011684 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 229.64 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -56.12 | -2.95 | -5.56 | -0.55
| Weight Savings, Therm |  |  |  |  | -16.6074
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 199.129 | 198.721 | 198.721 | 198.721 | 198.721
| Total | 3623155 | 3619557 | 3618604 | 3619478 | 3618839 |  | Total | 601.1008 | 600.6869 | 600.6661 | 600.688 | 600.619
Natural Gas Consumption (kBtu) | Total | 6409254 | 6422832 | 6410011 | 6410521 | 6409397
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.408 | 0.408
| Total |  | 3598 | 4551 | 3677 | 4316 |  | Total |  | 0.4139 | 0.4347 | 0.4128 | 0.4818
Natural Gas Savings (therms) | Total |  | -135.78 | -7.57 | -12.67 | -1.43
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0059 | 0.0267 | 0.0048 | 0.0738
| Indirect HVAC Savings, kWh |  | 2 | 956 | 82 | 721 |  | Weight Reduction, kW |  |  |  |  | 0.028084 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 556.08 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -135.78 | -7.57 | -12.67 | -1.43
| Weight Savings, Therm |  |  |  |  | -40.2886

## CZ08
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 198.129 | 197.959 | 197.959 | 197.959 | 197.959
| Total | 3615305 | 3613869 | 3613466 | 3613777 | 3613428 |  | Total | 531.2588 | 531.0848 | 531.065 | 531.0863 | 531.0622
Natural Gas Consumption (kBtu) | Total | 6507668 | 6513277 | 6507955 | 6508201 | 6507676
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1436 | 1839 | 1528 | 1877 |  | Total |  | 0.174 | 0.1938 | 0.1725 | 0.1966
Natural Gas Savings (therms) | Total |  | -56.09 | -2.87 | -5.33 | -0.08
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.004 | 0.0238 | 0.0025 | 0.0266
| Indirect HVAC Savings, kWh |  | -56 | 347 | 36 | 385 |  | Weight Reduction, kW |  |  |  |  | 0.016656 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 212.5 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -56.09 | -2.87 | -5.33 | -0.08
| Weight Savings, Therm |  |  |  |  | -16.4444
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 199.129 | 198.721 | 198.721 | 198.721 | 198.721
| Total | 3625262 | 3621662 | 3620728 | 3621580 | 3620780 |  | Total | 532.3424 | 531.9247 | 531.8778 | 531.9286 | 531.8707
Natural Gas Consumption (kBtu) | Total | 6498334 | 6511592 | 6499028 | 6499589 | 6498353
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.408 | 0.408
| Total |  | 3600 | 4534 | 3682 | 4482 |  | Total |  | 0.4177 | 0.4646 | 0.4138 | 0.4717
Natural Gas Savings (therms) | Total |  | -132.58 | -6.94 | -12.55 | -0.19
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0097 | 0.0566 | 0.0058 | 0.0637
| Indirect HVAC Savings, kWh |  | 4 | 939 | 87 | 887 |  | Weight Reduction, kW |  |  |  |  | 0.03973 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 583.26 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -132.58 | -6.94 | -12.55 | -0.19
| Weight Savings, Therm |  |  |  |  | -38.9296

## CZ09
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 198.129 | 197.959 | 197.959 | 197.959 | 197.959
| Total | 3666577 | 3665077 | 3664600 | 3665046 | 3664753 |  | Total | 595.2666 | 595.0934 | 594.9922 | 595.09 | 595.0667
Natural Gas Consumption (kBtu) | Total | 6608212 | 6613806 | 6608257 | 6608744 | 6608224
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1500 | 1977 | 1531 | 1824 |  | Total |  | 0.1732 | 0.2744 | 0.1766 | 0.1999
Natural Gas Savings (therms) | Total |  | -55.94 | -0.45 | -5.32 | -0.12
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0032 | 0.1044 | 0.0066 | 0.0299
| Indirect HVAC Savings, kWh |  | 8 | 485 | 39 | 332 |  | Weight Reduction, kW |  |  |  |  | 0.051452 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 276.86 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -55.94 | -0.45 | -5.32 | -0.12
| Weight Savings, Therm |  |  |  |  | -15.3958
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 199.129 | 198.721 | 198.721 | 198.721 | 198.721
| Total | 3676656 | 3673022 | 3671846 | 3672938 | 3672244 |  | Total | 596.4928 | 596.0623 | 595.8341 | 596.069 | 596.0134
Natural Gas Consumption (kBtu) | Total | 6599192 | 6612590 | 6599276 | 6600458 | 6599194
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.408 | 0.408
| Total |  | 3634 | 4810 | 3718 | 4412 |  | Total |  | 0.4305 | 0.6587 | 0.4238 | 0.4794
Natural Gas Savings (therms) | Total |  | -133.98 | -0.84 | -12.66 | -0.02
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0225 | 0.2507 | 0.0158 | 0.0714
| Indirect HVAC Savings, kWh |  | 38 | 1215 | 123 | 817 |  | Weight Reduction, kW |  |  |  |  | 0.12732 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 698.34 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -133.98 | -0.84 | -12.66 | -0.02
| Weight Savings, Therm |  |  |  |  | -36.7108

## CZ10
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 198.129 | 197.959 | 197.959 | 197.959 | 197.959
| Total | 3635386 | 3633876 | 3633435 | 3633855 | 3633524 |  | Total | 640.5819 | 640.3774 | 640.4111 | 640.41 | 640.3816
Natural Gas Consumption (kBtu) | Total | 6986787 | 6992352 | 6986773 | 6987318 | 6986794
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1510 | 1951 | 1531 | 1862 |  | Total |  | 0.2045 | 0.1708 | 0.1719 | 0.2003
Natural Gas Savings (therms) | Total |  | -55.65 | 0.14 | -5.31 | -0.07
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0345 | 0.0008 | 0.0019 | 0.0303
| Indirect HVAC Savings, kWh |  | 18 | 459 | 39 | 370 |  | Weight Reduction, kW |  |  |  |  | 0.015594 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 276.14 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -55.65 | 0.14 | -5.31 | -0.07
| Weight Savings, Therm |  |  |  |  | -15.0614
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 199.129 | 198.721 | 198.721 | 198.721 | 198.721
| Total | 3645407 | 3641770 | 3640732 | 3641718 | 3640944 |  | Total | 641.6817 | 641.1907 | 641.2712 | 641.2688 | 641.2008
Natural Gas Consumption (kBtu) | Total | 6977966 | 6991229 | 6977939 | 6979223 | 6977983
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.408 | 0.408
| Total |  | 3637 | 4675 | 3689 | 4463 |  | Total |  | 0.491 | 0.4105 | 0.4129 | 0.4809
Natural Gas Savings (therms) | Total |  | -132.63 | 0.27 | -12.57 | -0.17
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.083 | 0.0025 | 0.0049 | 0.0729
| Indirect HVAC Savings, kWh |  | 41 | 1080 | 94 | 868 |  | Weight Reduction, kW |  |  |  |  | 0.037798 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 649.14 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -132.63 | 0.27 | -12.57 | -0.17
| Weight Savings, Therm |  |  |  |  | -35.9128

## CZ11
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 182.7513 | 182.5813 | 182.5813 | 182.5813 | 182.5813
| Total | 3574046 | 3572543 | 3572126 | 3572516 | 3572146 |  | Total | 564.1481 | 563.9757 | 563.8858 | 563.9722 | 563.9513
Natural Gas Consumption (kBtu) | Total | 9058782 | 9064265 | 9058599 | 9059291 | 9058790
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1503 | 1920 | 1530 | 1900 |  | Total |  | 0.1724 | 0.2623 | 0.1759 | 0.1968
Natural Gas Savings (therms) | Total |  | -54.83 | 1.83 | -5.09 | -0.08
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0024 | 0.0923 | 0.0059 | 0.0268
| Indirect HVAC Savings, kWh |  | 11 | 428 | 38 | 408 |  | Weight Reduction, kW |  |  |  |  | 0.045458 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 268.78 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -54.83 | 1.83 | -5.09 | -0.08
| Weight Savings, Therm |  |  |  |  | -14.114
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 183.7513 | 183.3433 | 183.3433 | 183.3437 | 183.3437
| Total | 3584101 | 3580478 | 3579467 | 3580415 | 3579523 |  | Total | 565.3359 | 564.9216 | 564.7057 | 564.9138 | 564.8638
Natural Gas Consumption (kBtu) | Total | 9050146 | 9063407 | 9049680 | 9051370 | 9050165
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.4076 | 0.4076
| Total |  | 3623 | 4634 | 3686 | 4578 |  | Total |  | 0.4143 | 0.6302 | 0.4221 | 0.4721
Natural Gas Savings (therms) | Total |  | -132.61 | 4.66 | -12.24 | -0.19
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0063 | 0.2222 | 0.0145 | 0.0645
| Indirect HVAC Savings, kWh |  | 27 | 1039 | 91 | 983 |  | Weight Reduction, kW |  |  |  |  | 0.1096 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 650.92 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -132.61 | 4.66 | -12.24 | -0.19
| Weight Savings, Therm |  |  |  |  | -34.0282

## CZ12
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 182.7513 | 182.5813 | 182.5813 | 182.5813 | 182.5813
| Total | 3541744 | 3540246 | 3539881 | 3540216 | 3539853 |  | Total | 575.6338 | 575.4617 | 575.3643 | 575.4574 | 575.4339
Natural Gas Consumption (kBtu) | Total | 9183026 | 9188605 | 9182934 | 9183542 | 9183034
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1498 | 1863 | 1528 | 1891 |  | Total |  | 0.1721 | 0.2695 | 0.1764 | 0.1999
Natural Gas Savings (therms) | Total |  | -55.79 | 0.92 | -5.16 | -0.08
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0021 | 0.0995 | 0.0064 | 0.0299
| Indirect HVAC Savings, kWh |  | 6 | 371 | 36 | 399 |  | Weight Reduction, kW |  |  |  |  | 0.049084 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 241.5 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -55.79 | 0.92 | -5.16 | -0.08
| Weight Savings, Therm |  |  |  |  | -14.7542
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 183.7513 | 183.3433 | 183.3433 | 183.3437 | 183.3437
| Total | 3551704 | 3548095 | 3547211 | 3548022 | 3547152 |  | Total | 576.8361 | 576.423 | 576.1894 | 576.4129 | 576.3563
Natural Gas Consumption (kBtu) | Total | 9174158 | 9187596 | 9173934 | 9175382 | 9174176
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.4076 | 0.4076
| Total |  | 3609 | 4493 | 3682 | 4552 |  | Total |  | 0.4131 | 0.6467 | 0.4232 | 0.4798
Natural Gas Savings (therms) | Total |  | -134.38 | 2.24 | -12.24 | -0.18
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0051 | 0.2387 | 0.0156 | 0.0722
| Indirect HVAC Savings, kWh |  | 13 | 898 | 87 | 957 |  | Weight Reduction, kW |  |  |  |  | 0.11789 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 582.38 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -134.38 | 2.24 | -12.24 | -0.18
| Weight Savings, Therm |  |  |  |  | -35.5028

## CZ13
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 182.7513 | 182.5813 | 182.5813 | 182.5813 | 182.5813
| Total | 3691134 | 3689629 | 3689194 | 3689604 | 3689286 |  | Total | 567.0948 | 566.9212 | 566.8318 | 566.9187 | 566.8969
Natural Gas Consumption (kBtu) | Total | 8072747 | 8078261 | 8072654 | 8073263 | 8072754
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1505 | 1940 | 1530 | 1848 |  | Total |  | 0.1736 | 0.263 | 0.1761 | 0.1979
Natural Gas Savings (therms) | Total |  | -55.14 | 0.93 | -5.16 | -0.07
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0036 | 0.093 | 0.0061 | 0.0279
| Indirect HVAC Savings, kWh |  | 13 | 448 | 38 | 356 |  | Weight Reduction, kW |  |  |  |  | 0.046308 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 267.3 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -55.14 | 0.93 | -5.16 | -0.07
| Weight Savings, Therm |  |  |  |  | -14.579
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 183.7513 | 183.3433 | 183.3433 | 183.3437 | 183.3437
| Total | 3701191 | 3697564 | 3696513 | 3697504 | 3696738 |  | Total | 568.2862 | 567.8694 | 567.6557 | 567.864 | 567.8109
Natural Gas Consumption (kBtu) | Total | 8063949 | 8077250 | 8063711 | 8065191 | 8063969
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.4076 | 0.4076
| Total |  | 3627 | 4678 | 3687 | 4453 |  | Total |  | 0.4168 | 0.6305 | 0.4222 | 0.4753
Natural Gas Savings (therms) | Total |  | -133.01 | 2.38 | -12.42 | -0.2
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0088 | 0.2225 | 0.0146 | 0.0677
| Indirect HVAC Savings, kWh |  | 31 | 1083 | 92 | 858 |  | Weight Reduction, kW |  |  |  |  | 0.11103 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 645.56 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -133.01 | 2.38 | -12.42 | -0.2
| Weight Savings, Therm |  |  |  |  | -35.1134

## CZ14
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 182.7513 | 182.5813 | 182.5813 | 182.5813 | 182.5813
| Total | 3615751 | 3614241 | 3613788 | 3614219 | 3613867 |  | Total | 562.2111 | 562.0273 | 562.0142 | 562.0382 | 562.0161
Natural Gas Consumption (kBtu) | Total | 8958610 | 8963984 | 8957867 | 8959118 | 8958617
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1510 | 1963 | 1532 | 1884 |  | Total |  | 0.1838 | 0.1969 | 0.1729 | 0.195
Natural Gas Savings (therms) | Total |  | -53.74 | 7.43 | -5.08 | -0.07
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0138 | 0.0269 | 0.0029 | 0.025
| Indirect HVAC Savings, kWh |  | 18 | 471 | 40 | 392 |  | Weight Reduction, kW |  |  |  |  | 0.020234 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 285.7 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -53.74 | 7.43 | -5.08 | -0.07
| Weight Savings, Therm |  |  |  |  | -11.4754
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 183.7513 | 183.3433 | 183.3433 | 183.3437 | 183.3437
| Total | 3625811 | 3622171 | 3621154 | 3622120 | 3621306 |  | Total | 563.3112 | 562.8702 | 562.8389 | 562.8967 | 562.8432
Natural Gas Consumption (kBtu) | Total | 8950902 | 8963884 | 8949307 | 8952125 | 8950921
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.4076 | 0.4076
| Total |  | 3640 | 4657 | 3691 | 4505 |  | Total |  | 0.441 | 0.4723 | 0.4145 | 0.468
Natural Gas Savings (therms) | Total |  | -129.82 | 15.95 | -12.23 | -0.19
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.033 | 0.0643 | 0.0069 | 0.0604
| Indirect HVAC Savings, kWh |  | 44 | 1062 | 96 | 910 |  | Weight Reduction, kW |  |  |  |  | 0.048494 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 651 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -129.82 | 15.95 | -12.23 | -0.19
| Weight Savings, Therm |  |  |  |  | -28.5598

## CZ15
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 198.129 | 197.959 | 197.959 | 197.959 | 197.959
| Total | 3964923 | 3963383 | 3962860 | 3963355 | 3963056 |  | Total | 693.7497 | 693.5809 | 693.5799 | 693.5783 | 693.5524
Natural Gas Consumption (kBtu) | Total | 5593950 | 5599595 | 5594233 | 5594512 | 5593963
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1540 | 2063 | 1568 | 1867 |  | Total |  | 0.1688 | 0.1698 | 0.1714 | 0.1973
Natural Gas Savings (therms) | Total |  | -56.45 | -2.83 | -5.62 | -0.13
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0012 | -0.0002 | 0.0014 | 0.0273
| Indirect HVAC Savings, kWh |  | 48 | 571 | 76 | 375 |  | Weight Reduction, kW |  |  |  |  | 0.005232 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 336.42 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -56.45 | -2.83 | -5.62 | -0.13
| Weight Savings, Therm |  |  |  |  | -16.566
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 199.129 | 198.721 | 198.721 | 198.721 | 198.721
| Total | 3975154 | 3971424 | 3970287 | 3971433 | 3970779 |  | Total | 694.7904 | 694.3843 | 694.3824 | 694.3794 | 694.3171
Natural Gas Consumption (kBtu) | Total | 5584668 | 5597927 | 5585456 | 5585993 | 5584706
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.408 | 0.408
| Total |  | 3730 | 4867 | 3721 | 4375 |  | Total |  | 0.4061 | 0.408 | 0.411 | 0.4733
Natural Gas Savings (therms) | Total |  | -132.59 | -7.88 | -13.25 | -0.38
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | -0.0019 | 2.8422e-14 | 0.003 | 0.0653
| Indirect HVAC Savings, kWh |  | 134 | 1272 | 126 | 780 |  | Weight Reduction, kW |  |  |  |  | 0.012926 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 740.2 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -132.59 | -7.88 | -13.25 | -0.38
| Weight Savings, Therm |  |  |  |  | -39.449

## CZ16
ULT Freezer Range: 15 to <24 ft3 |  | Average ULT Freezer Direct Savings: 34%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1469900 | 1468408 | 1468408 | 1468408 | 1468408 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 182.7513 | 182.5813 | 182.5813 | 182.5813 | 182.5813
| Total | 3428977 | 3427477 | 3427113 | 3427442 | 3427027 |  | Total | 494.9296 | 494.7466 | 494.745 | 494.7562 | 494.7181
Natural Gas Consumption (kBtu) | Total | 13166487 | 13172138 | 13164748 | 13166972 | 13166493
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 1492 | 1492 | 1492 | 1492 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.17 | 0.17 | 0.17 | 0.17
| Total |  | 1500 | 1864 | 1535 | 1950 |  | Total |  | 0.183 | 0.1846 | 0.1734 | 0.2115
Natural Gas Savings (therms) | Total |  | -56.51 | 17.39 | -4.85 | -0.06
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.013 | 0.0146 | 0.0034 | 0.0415
| Indirect HVAC Savings, kWh |  | 8 | 372 | 43 | 458 |  | Weight Reduction, kW |  |  |  |  | 0.01822 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 255.08 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -56.51 | 17.39 | -4.85 | -0.06
| Weight Savings, Therm |  |  |  |  | -7.9828
ULT Freezer Range: 24-29 ft3 |  | Average ULT Freezer Direct Savings: 54%
Energy Usage | [6 empty] | Electric Demand
eQuest Output: |  | 0 | 1 | 2 | 3 | 4 | eQuest Output: |  | 0 | 1 | 2 | 3 | 4
| Base case: | -- | Base case | Base case | Base case | Base case |  | Base case: | -- | Base case | Base case | Base case | Base case
Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT | Climate Zone: CZ01 | Model Run: | Base case | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Electricity Consumption (kWh) | Misc Equip (inc. ULT freezers) | 1478661 | 1475065 | 1475066 | 1475066 | 1475066 | Electric Peak Demand (kW) | Misc Equip (inc. ULT freezers) | 183.7513 | 183.3433 | 183.3433 | 183.3437 | 183.3437
| Total | 3439058 | 3435446 | 3434539 | 3435360 | 3434337 |  | Total | 496.0338 | 495.5949 | 495.5916 | 495.6189 | 495.527
Natural Gas Consumption (kBtu) | Total | 13160129 | 13173575 | 13155915 | 13161265 | 13160148
Savings Estimate: | [6 empty] | Demand Reduction Estimate:
Electricity Savings (kWh) | Misc Equip (inc. ULT freezers) |  | 3596 | 3595 | 3595 | 3595 | Peak Demand Reduction (kW) | Misc Equip (inc. ULT freezers) |  | 0.408 | 0.408 | 0.4076 | 0.4076
| Total |  | 3612 | 4519 | 3698 | 4721 |  | Total |  | 0.4389 | 0.4422 | 0.4149 | 0.5068
Natural Gas Savings (therms) | Total |  | -134.46 | 42.14 | -11.36 | -0.19
|  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT |  |  |  | Lab ULT | Equip ULT | Corr ULT | Farm ULT
Weighting % in Location |  |  | 26% | 42% | 12% | 20% | Weighting % in Location |  |  | 26% | 42% | 12% | 20%
[8 empty] | Indirect HVAC kW reduction |  | 0.0309 | 0.0342 | 0.0073 | 0.0992
| Indirect HVAC Savings, kWh |  | 16 | 924 | 103 | 1126 |  | Weight Reduction, kW |  |  |  |  | 0.043114 | #DIV/0!
|  | Weight Savings, kWh |  |  |  | 629.8 | [7 empty] | #REF!
| Indirect HVAC Savings, Therm |  | -134.46 | 42.14 | -11.36 | -0.19
| Weight Savings, Therm |  |  |  |  | -18.662

## Sheet1
| Miscellaneous equipment end-use energy | Total end-use energy
| G4 | Q4
cz1 | AVERAGE(G5976:G5978,G6000:G6002,G6024:G6026) | AVERAGE(Q5976:Q5978,Q6000:Q6002,Q6024:Q6026)
cz2 | AVERAGE(G4536:G4538,G4560:G4562,G4584:G4586) | AVERAGE(Q4536:Q4538,Q4560:Q4562,Q4584:Q4586)
cz3 | AVERAGE(G4536:G4538,G4560:G4562,G4584:G4586)
cz4 | AVERAGE(G5856:G5858,G5880:G5882,G5904:G5906)
cz5 | AVERAGE(G6024:G6026,G6048:G6050,G6072:G6074)
cz6 | AVERAGE(G5856:G5858,G5880:G5882,G5904:G5906)
cz7 | AVERAGE(G5856:G5858,G5880:G5882,G5904:G5906)
cz8 | AVERAGE(G5856:G5858,G5880:G5882,G5904:G5906)
cz9 | AVERAGE(G5856:G5858,G5880:G5882,G5904:G5906)
cz10 | AVERAGE(G5856:G5858,G5880:G5882,G5904:G5906)
cz11 | AVERAGE(G4536:G4538,G4560:G4562,G4584:G4586)
cz12 | AVERAGE(G4536:G4538,G4560:G4562,G4584:G4586)
cz13 | AVERAGE(G4536:G4538,G4560:G4562,G4584:G4586)
cz14 | AVERAGE(G5712:G5714,G5736:G5738,G5760:G5762)
cz15 | AVERAGE(G5712:G5714,G5736:G5738,G5688:G5690)
cz16 | AVERAGE(G4536:G4538,G4560:G4562,G4584:G4586)